## Console Status
CONSOLE HEALTH REPORT
|  | SUBSYSTEM | SUBSYSTEM STATUS | Failures/Needed Actions
Date of Analysis: | July 24, 2025 | Host Hardware | Host Disk Failure Detected
Product Type: | Discovery MI DR | Hardware | NO
Software Build: | pet_columbia.179.H40_PN_CM64_G_GTLNI:9x6x4_lyso
System ID: | IL1062NM34 | Recon Queues | Recon Queue Problems Detected
Hospital Name: | Hadassah Medical Centers" | Software | NO
Console Type: | NIO64
HP Computer: | HP Z840 Workstation Serial Number: 2UA736272V | ASIR | Recon GPU Problems Detected
ASIR Exists: | Yes | Hardware | NO
Last Software Reboot: | 18 Minutes ago
Host GPU Core Temp: | 44 Celsius | Scan RAID | Scan RAID Disk Problem Detected | sg_map -i; sudo gre-raid -q -v ¦ grep 'Device Id';
ASIR Temp: | 36 | Hardware | NO | smartctl --all --device=megaraid, <device Id> /dev/sdc1;
HP BIOS Level: | BIOS Revision: 1.62 |  |  | Execute command on devices to determine whether device needs replacement
Manufactoring Date: | 2017 week 36 | Image DB | Image Database Errors Detected
Warranty End Date: | 2020 week 48 | Software | NO
SUSE Version: | SUSE Linux Enterprise Server 11 (x86_64)
PNF Status: | ON | Scan DB | Scan Database Errors Detected
Software Installation I/O Error: | None | Software | NO
Eth2: | 192.168.40.34
eth2 Duplex: Full
eth2 Speed: 1000Mb/s
FOR ANY SOFTWARE RELATED ISSUES, PLEASE DO LFC OR UPDATE SOFTWARE OR RESET IMAGE DB (DON'T REPLACE HARDWARE)

## High Voltage Components
Property of GE Healthcare Systems | Intelligent Diagnostic Solutions | High Voltage Evaluation Report The 6 categories below are standalone and will report suggestions if issues are found. Use the Testing Tabs to identify failing component in each category. | Possible Faulty Part | T/S Tab
System Identification Number | IL1062NM34 | No Filament Aging Problems Found
Product Type | Discovery MI DR | No Rotation Problems Found
Software-Hardware Configuration | pet_columbia.179.H40_PN_CM64_G_GTLNI 9x6x4_lyso | No Filament Problems Found
High Voltage Generator Type | Jedi HV w/ CjM40 Liquid Bearings Tube | No High Voltage Problems Found
Hospital Name | Hadassah Medical Centers | No Power Problems Found
Date Analysis Performed | 2025-7-24 5:8:41 UTC+00:00 | No Hardware Problems Found
Note: No issue found may exist for non-updated Jedi Log. Verify gesyslog tab for errors
Reg & Spits Analysis | Existing Tube Information
Anode Spits | Insert Serial Number : 451141TU6 | IMPORTANT
0 | Insert Model Number : 5401074 | TubeWatch Predictive triggers may not appear in HVBAT as they are primarily derived from model-based predictions.
Cathode Spits | Casing Serial Number : 106530ty9
0 | Casing Model Number : 2137130
Insert Spits | Install Date : Thu Jul 24 07:37:53 2025
0 | Last Scan Taken On : no scans taken
Reg (Exp) Spits | Total Patient Exams : 0
0 | Total mAs : 0
Reg (Rise) Spits
0 | Tube Manufacturer: GE
Type in findings, testing, parts used, and solution here:

## CT Detectors Status
Detector BAT |  | Detector Details |  |  |  | Filter on "Latest Error TimeStamp" column for most recent errors
|  | Slot | FRU Part Number | Detected Issue | Recommended Action | Latest Error TimeStamp | Module Change Timestamp
Date of Analysis | Thursday, July 24, 2025 5:11 AM | Module-1 | 5481498-9 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
SystemId | IL1062NM34 | Module-2 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
Hospital Name | Hadassah Medical Centers | Module-3 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
Det Type | CHDAS_MERC_64 | Module-4 | 5481498-11 | No Problem Reported.
SW Build | pet_columbia.179.H40_PN_CM64_G_GTLNI:9x6x4_lyso | Module-5 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
Access Level | Unknown | Module-6 | 5481498-11 | Passes Exclusion Test OPEN channel criteria. |  | Thu Jul 24 07:07:00 2025
Det Serial Number | See Module Hart Info Tab | Module-7 | 5481498-10 | No Problem Reported.
LPP | Not Installed | Module-8 | 5481498-10 | No Problem Reported.
Total Parse Time: | 0.079 Seconds | Module-9 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
| Other Details | Module-10 | 5481498-10 | No Problem Reported. |  |  | Module changed on: Thursday, April 17, 2025 11:48:18 AM
| Plenum or Rail temperature is low.If issue persists, Confirm room temperature is in proper range and confirm that gantry heater is operational. Thu Jul 24 07:44:32 2025 | Module-11 | 5481498-10 | No Problem Reported.
|  | Module-12 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-13 | 5481498-9 | No Problem Reported.
|  | Module-14 | 5481498-9 | No Problem Reported.
|  | Module-15 | 5481498-9 | No Problem Reported.
|  | Module-16 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-17 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-18 | 5481498-10 | No Problem Reported.
|  | Module-19 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-20 | 5481498-10 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-21 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-22 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-23 | 5481498-11 | No Problem Reported.
|  | Module-24 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-25 | 5481498-11 | Passes Exclusion Test SHORT channel criteria. |  | Thu Jul 24 07:07:00 2025
|  | Module-26 | 5481498-11 | No Problem Reported.
|  | Module-27 | 5481499-4 | No Problem Reported.

## Tube Hist
Tube | Date Installed | Tube #1
1 | Thu Jul 24 07:37:53 2025 | Insert Serial Number | 451141TU6
2 | Sat Oct 14 01:09:45 2017 | Insert Model Number | 5401074
3 | Fri Oct 13 03:32:45 2017 | Casing Serial Number | 106530ty9
|  | Casing Model Number | 2137130
|  | Install Date | Thu Jul 24 07:37:53 2025
|  | Last Scan Taken On | no scans taken
|  | Total Patient Exams | 0
|  | Total mAs | 0
Tube #2
|  | Insert Serial Number | 209314GI4
|  | Insert Model Number | 5401074
|  | Casing Serial Number | 100516TY4
|  | Casing Model Number | 2137130
|  | Install Date | Sat Oct 14 01:09:45 2017
|  | Last Scan Taken On | Thu 24 Jul 2025 07:19:10 AM EEST
|  | Total Patient Exams | 31175
|  | Total mAs | 159173558
Tube #3
|  | Insert Serial Number | 209314GI4
|  | Insert Model Number | 5401074
|  | Casing Serial Number | 100516TY4
|  | Casing Model Number | 2137130
|  | Install Date | Fri Oct 13 03:32:45 2017
|  | Last Scan Taken On | Fri 13 Oct 2017 10:31:25 PM CDT
|  | Total Patient Exams | 0
|  | Total mAs | 17820

## Option List
----Option List----
Product Software Option Keys:
PET Base Permanent
PET Dynamic Review Permanent
PET List Mode Permanent
TOF Permanent
VolumeViewer Permanent
VolumeViewerPet Permanent
Smart Prep Permanent
Connect Pro Permanent
Smart Speed Permanent
Patient-64-slice Permanent
Direct-MPR Permanent
Data Export Permanent
CopyComposer Permanent
3000 Image Series Permanent
AutomA Permanent
WideView Permanent
PET Diagnostic CT Permanent
PET 3D Iterative Recon Permanent
Sub-0.4-Second-Scan Permanent
PET Gating Permanent
PET Cardiac Permanent
PET Dynamic Permanent
SmartScore Pro Permanent
EKG Viewer Permanent
Edge Plus Kernel Permanent
Exam Split Permanent
Rad_Rx Permanent
Organ Dose Modulation Permanent
ASIR Permanent
AWE connection Permanent
100kVA Permanent
Cine-120-Sec Permanent
Dynamic Z-axis Tracking Permanent
Helical-120-Sec Permanent
High Pitch Helical - 1.53 Permanent
Ultra Kernel Permanent
Image Check Permanent
Dynamic Transition Permanent
0_4Speed Permanent
MotionFreeApps Permanent
PET Q.Core Power Permanent
NoiseReductionFilter Permanent
ACQC Permanent
PET Cardiac Shift Permanent
CardIQSnapShot-Cine Permanent
CardIQ SnapShot Permanent
Enhanced Xtream Injector Permanent
Liquid-0.35-Routine Permanent
PET Cardiac Review Permanent
PET Q.AC Permanent
PET Q.Freeze Permanent
PET Q.Static Permanent
PET Regularized Recon Permanent
Overlapped Recon - Axial Permanent
SharpIR Permanent